ESA609 Firmware Update Instructions

Firmware version 1.01

1. Download and install the latest Ansur Executive software from the Fluke Biomedical website:

http://www.flukebiomedical.com/Biomedical/usen/Support/software/Ansur-Software.htm

2. Download the ESA609 firmware v1.01 from the ESA609 product page:

https://www.flukebiomedical.com/products/biomedical-test-equipment/electrical-safety-analyzers/esa609-electrical-safety-analyzer

3. Connect ESA609 to PC via USB (keep ESA609 unplugged/powered off).

4. Open the ‘Update Firmware’ application shortcut (under the ‘Fluke’ start menu group) to launch the ‘Ansur Firmware Upload’ utility.

5. Select the ESA609 product from the devices list.

- Click the ‘Scan ports…’ link.

- Select the COM port corresponding to the connected ESA609 from the ports dropdown list.

- Click the ‘Next >’ button.

3. Browse to and select the downloaded firmware file ‘ESA609fw_v1.01.01.hex’ when prompted.

- Click the ‘Next >’ button

4. Follow the on-screen instructions for initiating the firmware update:

- Hold F1 and F3 keys simultaneously on the front panel of the ESA609 while connecting the inlet power cord.

Note: Device will power on displaying a blank screen.

- Press the ‘Upload >’ button.

- Restart the ESA609 when prompted by unplugging and reinserting the inlet power cord.