ESA620 Update Instructions

Firmware version 2.12

Instructions for updating your ESA620:

- If necessary, download and install Ansur Executive 2.9.7 (or later) and Ansur ESA620 Plug-In 1.1.12 (or later) from http://www.flukebiomedical.com/Biomedical/usen/Support/software/Ansur-Software.htm

- Extract the files from this zip archive.

- Run the Update Firmware tool via the Start Menu > All Programs > Fluke > Update Firmware.

- Select ESA620 and select the appropriate COM port and click Next.

- Click the first “…” button associated with UI Firmware file and find “ESA620_UIFw_2.12.02.hex” which was extracted from the zip archive.

- Click the second “…” button associated with DSP Firmware file and find “ESA620_DspFw_2.12.02.x00” which was extracted from the zip archive.

- Click Next.

- Plug the ESA620 into your PC and power it on.

- Click Upload.

- After the upload finishes, your unit should boot normally. No calibration is required.