Impulse 7000 Update Instructions

Firmware version 2.12

Instructions for updating your Impulse 6000 and 7000:

- If necessary, download and install Ansur Executive 3.1.3 (or later) and Ansur Impulse 7000 Plug-In 1.2.5 (or later) from http://www.flukebiomedical.com/Biomedical/usen/Support/software/Ansur-Software.htm

- Extract the files from this zip archive.

- Run the Update Firmware tool via the Start Menu > All Programs > Fluke > Update Firmware.

- Select Impulse 6000/7000 and select the appropriate COM port and click Next.

- Click the first “…” button associated with UI Firmware file and find “Imp67UIFw_2.12.hex” which was extracted from the zip archive.

- Click the second “…” button associated with DSP Firmware file and find “Imp67DspFw_2.12.x00” which was extracted from the zip archive.

- Click Next.

- Plug the Impulse 6000 or 7000 into your PC.  Press SETUP and F4 on the front panel of the Impulse while turning on the power.

- Click Upload.

- When the UI firmware is uploaded successfully, the Impulse will reboot and load the DSP firmware.  After the upload finishes, your unit should boot normally. No calibration is required.