ProSim 8/8P Firmware Update Instructions

Firmware version 2.10

This firmware improvement features:

- Corrected bug that when a unit is set to have a heart rate of 360 BPM and then the user swaps to SPO2 it appears that the Prosim8 lost its connection to the SPOT unit. Now it’s correctly displaying a warning to the user explaining that the heart rate is too high.

- Corrected bug that if the Masimo Rainbow SpO2 type is set and the system is power cycled the connect window will appear on the screen before the unit is finished powering on. Now the correct window is correctly only appearing when entering the SPO2 menu.

- Fixed translated texts on the main NIBP menu, the 5th line read "wave" instead of which should be the translation of "Envelope Shift". Also fixed the text that if the user selected the last field (5th line), the resulting sub-menu displays the translation for "Heart Rate" when it should instead be the translation for "Envelope Shift".

- Fixed Chinese, Japanese, French and German translations.

Detailed firmware upgrade instructions are available in the ProSim 8 Ansur plug-in user manual. The update instructions apply for both ProSim 8 and ProSim 8P. The following is a brief instruction:

1. Download both the latest Ansur Executive and Ansur ProSim 8 Plug-in (including ProSim 8 Mini Plug-in) from http://www.flukebiomedical.com website.

2. Open the Mini Plug-In.

3. Click the Update Firmware tab.

4. Click next to the field of the firmware you want to upload. Upload the main unit processor and user interface firmware files.  If you have an SpO2 accessory, connect it to the ProSim 8/8P and also upload the SPOT firmware file.

5. Click the Next button, follow the instructions on the screen

6. Click the Upload button to send the file(s) to ProSim 8/8P.

Please contact your regional sales managers or Fluke Biomedical Technical Support at

techservices@flukebiomedical.com if you have any questions or concerns about this update.